ICAD IN SITU ANALYSERS

PATENTED, DIRECT NO2 AND / or HONO DETECTION - NITROGENMONOXIDE (NO) MEASUREMENT VIA O3 TITRATION CONVERTER - HIGH PRECISION – EASY OPERATION - CALIBRATION NOt REQUIRED

NO2/NO/HONO measurement - ppt range - high time resolution - low power consumption

High Sensitivity, LARGE dynamic range & Mobility

The ICAD features typical advantages of high accuracy, instrumental stability, long maintenance intervals and low consumables. Further, the high dynamic range allows measurements from high polluted conditions e.g., at high traffic roads or industrial monitoring to very low concentrations in clean environments. If even ultra-low NOx concentrations down to 15 ppt are of interest, the special ICAD high-grade versions “*L” are the perfect tool. The measurement can easily be controlled with a tablet connected to the ICAD via WiFi.

WORK SpaCE NO2 / NOx Monitoring or INDUSTRIAL monitoring

The high mobility, rugged design, and low maintenance effort make ICAD instrument the ideal instrumentation for reliable monitoring NOx levels in workplace environments such as mines, constructions site or industrial production places. Further, ICAD instruments can be applied to measure and monitor the NOx emissions of industrial machines. Multiple data interfaces enable optional integration of ICAD instruments with industrial processes.

Mobile measurements

The low power consumption, compact size, moderate weight and insensitivity to vibrations allow easily mobile measurements at different locations. The short set-up and warm-up time gives a lot of flexibility. Customized, ICAD versions for applications on drones are also available.

mobile applications – On-road real driving emissions

ICAD for emissions “*E” (equipped with additional internal CO2 sensor) allow on-road real driving emissions via the so-called Plume Chasing method. A vehicle with the ICAD, follows target vehicles to measure the gases in the diluted plume. Within seconds the system derives the specific NOx emission signature from the vehicle. In short times, high emitting vehicles due to defects or exhaust manipulations are identified. The tool allows authorities to enhance inspection efficiency or researchers to perform a vehicle emission screening.

flexible deployment – EXTENDED application

The fast response time of the ICAD instrument enables monitoring and assignment of NOx emissions from water vehicles by placing the instrument near water ways (e.g., at bridges or at the shoreline). With the ICAD emission configuration, NOx emission factors (mg NOx/kWh) are similar derived like for the real driving application. The calculation is independent from dilution of exhaust gases and thus, influence of meteorology.

No Drifts, no cross-interferences

The selective detection of NO2, NOx or HONO, high measurement sensitivity and negligible drift makes the ICAD perfectly suitable for stationary air quality monitoring as well as scientific studies of chemical processes, e.g., atmospheric studies or simulation experiments. Further, the absence or cross-interferences to other gas species, enables measurements also of complex mixture (e.g., containing high concentration of CO2, N2O, H2O or hydrocarbons).

Technique And Features

|  | - Direct measurement by differential absorption spectroscopy - High sensitivity and accuracy by use of ICAD algorithm (insensitive to intensity variations and aging of light sources) - No cross-interference, spectral separation from other gases like water vapour (H2O), Glyoxal (C2H2O2), oxygen (O2) - NO2 / HONO calibration gases not required; no pre-drying of sample air required - High dynamic detection range of low ppt to ppm - Fast response time of 1 seconds - Optional internal CO2 extension for combustion emission measurements (e.g., vehicle on-road measurement, Plume Chasing) - Remote access with any Wi-Fi device. GUI for easy operation. Multiple data interfaces. - Graphical system health table for easy and fast on-board diagnosis. |
| --- | --- |

ICAD model overview

|  | NO2-210 | NOx-210D | NOx-210DE | NO2-210L | NOx-210DL | HONO/NO2-210L |
| --- | --- | --- | --- | --- | --- | --- |
| Detectable gases | NO2 | NO2 / NO | NO2 / NO / CO2 | NO2 | NO2 / NO | HONO / NO2 |
| Range | 5 ppm | 5 / 5 ppm | 5 / 5 / 2000 ppm | 2 ppm | 2 ppm / 2 ppm | 0.5 / 2 ppm |
| Limit of detection at 2s, 30s, 300s in ppt | 350,100,30 ppt | 350,100,30 ppt | 350,100,30 ppt CO2: 4 ppm | 200,50,15 ppt | 200,50,15 ppt | HONO: 500,120,40 ppt NO2: 600,150,50 ppt |
| Precision (1σ) at 2s, 30s, 300s in ppt | 175,50,15 ppt | 175,50,15 ppt | 175,50,15 ppt CO2: 2 ppm | 100,25,8 ppt | 100,25,8 ppt | HONO: 250,60,20 ppt NO2: 300,75,25 ppt |
| Available also as mobile “M”-version |  |  |  |  |  |  |

Icad SPECIFICATIONS

| Detection of NO2, HONO | Direct spectroscopic measurement | Weight 19” Rack | < 12 kg (depending on config) |
| --- | --- | --- | --- |
| Detection of NO (NOx) | By conversion to NO2 | Weight “M” version | < 10 kg (depending on config) |
| Response Time (10% to 90%)*2 | 2s at 1 l/min or 1s at 2 l/min | Power consumption | Less than 40 W at 12 V (typ.) |
| Zero drift | Less than 0.1 ppb/month*3 | Start-up time | Less than 1 min (typ.) |
| Sample flow | 1 to 2 l/min | Temp. range of operation | -10 to +25°C (+40°C with cooling option) |
| Time resolution | 1s to 60 s temporal averaging | Temperature sensitivity | Less than 0.01 ppb/°C |
| Calibration for NO2 / HONO | Via spectroscopic data, gas-free*4 | Cross sensitivity | No significant cross sensitivity*5 |
| Path length characterization | Gas free ICOM method, Helium (optional, every 1 to 2 years) | Mech. stability | Insensitive to vibrations |
|  |  | Consumable gases | No gases needed for operation |
|  |  | Other detectable gases | Glyoxal*6 (respected by spectral analysis) |
| Housing Size 19” Rack | 43.8 x 13.3 (3HU) x 43.5 cm3 (WHD) | Processing unit | Internal embedded PC (WIN10) |
| Housing Size 19” Rack “L” | 43.8 x 13.3 (3HU) x 66.5 cm3 (WHD) | Data communication | LAN/WiFi/RS232/M2M/OPCUA; Bayern-Hessen Protocol; Voltage/Current Output |
| Housing Size “M” | 40.0 x 13.3 (3HU) x 30.0 cm3 (WHD) |  |  |

[*1] Custom specifications with different measurement range are possible. By reducing the measurement range better precision and LOD can be achieved. [*2] Response: Different measurement cell types are available, allowing different response times. [*3] Upper limit. Drift is negligible due to regularly automated reference measurements. [*4] Literature absorption data for target gas is used for gas quantification. [*5] No significant spectroscopic cross sensitivity found for: water, ozone, Glyoxal, Carbon Oxides, Methane, Formaldehyde, Hydrogen, Sulphide, Sulphur Dioxide, Chlorine, Chlorine Dioxide, Hydrogen Cyanide, Hydrogen Chloride, Phosphine, Hydrogen, Ammonia, Acetylene, Nitromethane, Ethylene, Ethanol, Methyl Mercaptan, Ethyl Mercaptan. [*6] For NO2-210 and NOx-210 models.

ADVANTAGES

| Benefits | Innovation |
| --- | --- |
| High measurement accuracy | - Series 210 Improvement: Enhanced accuracy at higher concentrations - Series 210 Improvement: Enhanced time resolution down to 1s - Direct spectroscopic gas measurement - High sensitivity, low measurement error - Fast measurement response - No zero-point or calibration drift, 100% reproducibility, no interferences - No sample pre-dryer needed |
| Simple and low costs operation | - NO2 / HONO calibration gas not required - Parallel NO measurement (with ozone titration converter) - No consumable gases needed - Robust setup, long lifetime |
| Flexible application | - Series 210 Improvement: System Health GUI for fast on-board diagnosis - Series 210 Improvement: Easy adjustment of span calibration in GUI - High stability (not sensitive to shocks, vibration, temperature) - Compact design, mobile application - Low power consumption and 12 V operation - Data Interfaces: WiFi, LAN, Machine2Machine, RS232, Analogue Volt./Cur. - Internal memory for up to 2 years of data |

Patents:  DE102015000423; EP3329251; US15/748,923; China ZL201680057099.6