ICAD NO2-NOx-NO Analyzer SERIES 210DL

Patented, fast, accurate and direct Nitrogen Dioxide Detection

| ICAD-NOx-210DL series featuring 19" rack housing and OLED display. | WI-FI LOW MAINTENANCE LOW POWER CONSUMPTION DIRECT NO2 DETECTION - Iterative Cavity DOAS (ICAD) spectroscopy - High resolved spectroscopic NO2 detection from 430 to 480 nm using fingerprint of NO2 spectrum - No observed interferences, e.g. to water, glyoxal, Ozone - No sample drying, e.g. Nafion tubing, required NOx MEASUREMENTS BY INTERNAL CONVERSION - Patented, NOx-free ozone generator for NO conversion - Dual measurement cells for parallel NO2 and NO detection LONG-TERM STABILITY AND ROBUSTNESS - Insensitive to temperature and vibrations - Independent from absolute light source intensity - No zero-point drift, integrated zero measurement EASY SET-UP AND OPERATION, LOW MAINTENANCE - No gas calibration required, minimal maintenance effort - Controllable via Wi-Fi devices (PC, tablet, smartphone) - Multiple data communication protocols - 19” rack - Sustainable: Long life time, low power consumption SERIE 210 NEW FEATURES - System health user interface and on-board diagnosis - Easy adjustment of span calibration in user interface - Enhanced time resolution of 1 second - Improved accuracy at high concentrations - Improved long term stability |
| --- | --- |
| 5600 s 1.7 ppt Time scaled modified Allan deviation for NO2 precision. Note: Detection limits are defined as twice the modified Allen deviation values. |  |
| Patents: EP3329251; D502016003239.4; GB3329251; FR3329251; US10473583B2; ZL201680057099.6; DE102015000423 (pending); EP325928; D502016014735.3; GB3325928; FR3325928 |  |

| ICAD Product webpage | Contact Airyx |
| --- | --- |

| GRAPHICAL USER INTERFACE AND ON-BOARD DIAGNOSIS |  | MEASUREMENT PRINCIPLE AND GAS FLOW |
| --- | --- | --- |
| The Graphical User Interface provides easy operation and quick access to advanced functions and data time series. The ICAD can be accessed and controlled by any Wi-Fi device (PC, tablet, smartphone) without additional software/apps required. A comprehensive system health table shows the health status of critical sensor data (e.g. gas flow, supply voltage, temperature) for easy and fast on-board diagnosis. |  | ICAD instruments are based on optical resonators which provide here long absorption paths of up to 4 km for precise NO2 / NOx detection especially at low concentrations. Gas concentrations are derived by patented Iterative Cavity DOAS (ICAD) spectroscopy between ~430 to 480nm. Below, the measurement scheme incl. additional NO to NO2 converter for NOx measurement and optional CO2 sensor is shown. |

INSTRUMENT Specifications

| Measurement range1 | 0 – 2000 ppb |  | Other detectable gases | Glyoxal, CO2 (optional NDIR sensor) |  |
| --- | --- | --- | --- | --- | --- |
| Time resolution | 1 sec minimum temporal averaging |  | Mechanical stability | Insensitive to vibrations |  |
| Limit of detection1 (2σ) | 0.03 ppb at 60 sec |  | Power consumption | Less than 40 W at 12 V (typ.) |  |
| Precision (1σ) | 0.12 ppb at 1 sec |  | Processing unit | Internal embedded PC, with data analysis and measurement software |  |
|  | 0.02 ppb at 60 sec |  |  |  |  |
| Detection of NO2 | Direct ICAD spectroscopy |  | Temp. range of operation | -10 to +40°C |  |
| Detection of NOx / NO | By conversion to NO2 |  | Temperature sensitivity | Less than 0.01 ppb/°C |  |
| Response time (10% to 90%) | <3 sec at 2 l/min (flow adjustable) |  | Cross sensitivity | No significant cross sensitivity5 |  |
| Zero Drift | Less than 0.1 ppb/month2 |  | Weight | Less than 12 kg (depending on config) |  |
| Sample flow | 1 to 2 l/min |  | Consumable gases | No gases needed for operation |  |
| Path length characterization | Helium (every 1 to 2 years, optional) or NO2 reference measurement or Gas free ICOM3 method |  | Dimensions W x H x D | 19” rack housing | 43.8 x 13.3 (3HU) x 66.5 cm |
| Calibration | NO2 calibration gas not needed4; NO gas for converter calibration above 1000 ppb (~every 12 months) |  | Data communication | LAN/Wi-Fi/RS232/M2M/OPCUA; Bayern-Hessen Protocol; Volt./Current Output |  |
|  |  |  | Start-up time | Less than 1 min (typ.) |  |
| Model variations | ICAD-NO2-210L ICAD-NOx-210DL | (NO2 only) (NO + NO2) |  |  |  |

[1] Custom configurations / measurement ranges available. By reducing the measurement range better precision and LOD can be achieved.

[2] Upper limit. Regularly automated zero measurement reduce measurement noise and guarantee also a negligible zero drift.

[3] ICOM (upcoming with software update): Integrated Calibration by means of optical Modulation. Patented gas free calibration method allows consumable free, automatic path length monitoring and calibration of ICAD measurement cells. Patents: EP325928; D502016014735.3; GB3325928; FR3325928

[4] Literature absorption data for NO2 is used for gas quantification.

[5] No significant spectroscopic cross sensitivity to: Carbon oxides, Formaldehyde, Ozone, Methane, Hydrogen, Sulphide, Sulphur dioxide, Chlorine, Chlorine dioxide, Hydrogen cyanide, Hydrogen chloride, Phosphine, Hydrogen, Ammonia, Acetylene, Nitromethane, Ethylene, Ethanol, Methyl mercaptan, Ethyl mercaptan.